## Slide 1: Vertical cryostat
Final Design
Cryogenic and mechanical configurations
10 th May 2016
ACS – P. Bujard, J.P. Thermeau

## Slide 2: Outline
Cryogenic Studies – JPT
- Saturated or pressurized baths,
Tests of superconducting cavities,
- Cryostat used as vacuum vessel,
- Cryostat in liquid helium operations,
Tests of superconducting magnets,
  Cryostat in pressurized liquid helium operations.
Specifications to finalize.
Mechanical Studies – PB
- General view of the vertical cryostat facilities,
Valve box,
Cryostat,
Lambda Insert,
Helium gas heater,
Latest documents.

## Slide 3: Saturated or pressurized baths
Two possibilities to operate in superfluid helium:
with saturated liquid helium
with pressurized liquid helium
Power density in turbulent flow:

## Slide 4: Saturated liquid helium bath
Disadvantages:
Small temperature margin, the temperature of the cavity walls are closed to the saturated temperature, T = f( x g x h),
h = 0.2 m, p  3 mbar  T  40 mK
High quantity of liquid helium operates in sub-atmospheric condition.
Advantages:
High level of heat transfer in superfluid helium (Kapitza thermal resistance limitation),
High level of temperature stability by pressure regulation,
Simple mechanical design.
150 mm
488 mm
609 mm
 line
 point
GAS
SUPERFLUID
(He II)
Bi-phase pipe
Top of the cavity
Bottom of the cavity
Margin: 75 mK
Margin: 25 mK

## Slide 5: Pressurized liquid helium bath
Disadvantages:
Complex mechanical design,
Complex cryogenic design (pressurized helium bath cooled with a heat exchanger and saturated superfluid helium, …).
Advantages:
High level of heat transfer in superfluid helium
(Kapitza thermal resistance limitation),
Temperature margin of  0.3 K,
Better electrical isolation than at low pressure
Main Helium baths operate at a pressure higher than the atmospheric pressure (no plug with frozen air).

## Slide 6: Tests of superconducting cavities - Cryostat as vacuum vessel
Thermal losses of the Valve Box and Cryostat:
LN2 circuits: 20 + 50 = 70 W
4 K circuit (VB + line): 2 + 2 = 4 W
2 K circuit (Cryostat + line): 1 + 2 = 3 W
Specifications:
Operating pressure: 5 mbar to 1.5 bar,
Temperature range: 4.5 K to 1.6 K,
Cryogenic power: 90 W at 1.9 K.
Cryostat in vacuum condition
This configuration allows similar operations which can be made with the horizontal cryostat.

## Slide 7: Tests of superconducting cavities Cryostat in liquid helium operations
Thermal losses of the Valve Box and Cryostat:
LN2 circuits: 20 + 50 + 100 = 170 W
4 K circuit (VB): 2 W
2 K circuit (Cryostat): 5 W
Specifications:
Operating pressure: 5 mbar to 1.5 bar,
Temperature range: 1.6 K to 4.5 K,
Cryogenic power: 90 W at 1.9 K.
Neck thermalization
LN2, 85 K: 100 W,
Supercritical helium, 20 K: 0.1 g/s,
4 K thermal losses:
2 W.

## Slide 8: Tests of superconducting magnets Cryostat in pressurized liquid helium operations
Thermal losses of the Valve Box and Cryostat:
LN2 circuits: 20 + 50 + 100 = 170 W
4 K circuit (VB): 2 W
2 K circuit (Cryostat + Lambda plate): 5 + 15 = 20 W
Specifications:
Operating pressure: 1 bar to 5 bar,
Temperature range: 4.5 K to 1.6 K,
Cryogenic power: 90 W at 1.9 K.
Neck thermalization
LN2, 85 K: 100 W,
Supercritical helium,
20 K: 0.1 g/s,
4 K thermal losses: 2 W.
HX681 (1 m²)
Saturated helium: 1.8 K,
Power (design): 120 W,
Bath temperature: 1.86 K
(T maxi: 60 mK)

## Slide 9: Cryogenic power limits - Heat Exchanger HX681
Low pressure circuits specifications:
Operating pressure: 16 mbar (10 mbar to 1 bara),
Operating temperature: 1.8 K (1.6 K to 4.5 K),
Cryogenic power (design): 120 W at 1.8 K,
Inner pipe surface: 1m².
Atmospheric pressure circuits specifications:
Operating pressure: 1.2 bara (1 to 5 bara),
Operating temperature: 1.85 K (1.6 K to 4.5 K),
Cryogenic power (design): 120 W at 1.85 K.
Low pressure inside the heat exchanger
Atmospheric pressure outside the heat exchanger
Pressurized helium bath
He II
He II
Bath
He II
Bath
T(K)
1.86 K
1.8 K
Copper pipe
Kapitza T
Temperature of the pressurized helium bath
Temperature inside HX681
Power available for the magnet tests:
90 W (1.85 K) - Lambda leak - thermal losses  70 W

## Slide 11: Pressure vessel
Specifications:
Maximum allowable operating pressure:  4 barg,
Hydraulic test pressure: 7.15 bara,
The mechanical calculations refer to EN13445 and EN13458 norms
Dimensions:
height ~ 4.5 m
diameters: 1.1 to 1.25 m
Volume: 4.8 m3
Pressure vessel of category III, MAWP  4 barg
4.8 m3
Norms:
EN13458: Cryogenic vessels
EN13455: Unfired Pressure vessels

## Slide 12: Specifications to finalize
Pressure vessel specifications:
Maximum allowable operating pressure:  4 barg,
Hydraulic test pressure: 7.15 bara,
The first mechanical calculations refer to EN13445 and EN13458 norms
Define the norms for the manufacturing and the final calculations.
HX681, Low pressure circuits specifications:
Operating pressure: 16 mbar (10 mbar to 1 bara),
Operating temperature: 1.8 K (1.6 K to 4.5 K),
Inner pipe surface: 1m² (minimum value).
Define the needs for the cryogenic power at 1.8 K.
Magnet test specifications (next meeting at CERN):
Cool down sequence: thermal gradient, speed, …
Operating temperatures: 1.8 K to 80 K,
Define all sequences needed for the tests (electrical tests at 80 K, superconducting measurements at 4.5 K and 1.8 K, …),
Define the specifications for the cryogenic sequences (cryogenic and electrical power, instrumentation, …
Impact on the existing facilities and necessary developments:
Helium liquefier and its LHe storage
Cryogenic power at 1.85 K and the Helium pumping system
The magnet quenchs and the helium recovery system,
New cryostat with its own control system, the links with the horizontal cryostat and the other equipments.

## Slide 13: FREIA Hall
Gersemi hole:
3.3x3x1m
+
Ø2x4m

## Slide 14: Global scheme
Hnoss Reheater
Valve Box
Liquefier Dewar
Additional bottle
Gersemi cryostat
Liquefier
Filling station

## Slide 15: Valve Box
4K pot
80K shield
Body: Ø1000x2375
Overall height: ~2700
Weight: 1075kg
Ø1100
GHe pumping-port
Supercritical GHe Pre-cooling Hx
LHe sub-cooling Hx

## Slide 16: Vacuum Vessel
Multi line port
80K shield
2mm thick
ID: ~1650
inner height: 4200
Body:Ø1957x4330
Weight:2188kg
main vessel
Ø1800
8mm thick
ID : ~1784
inner height : ~4300
vacuum pumping-port
Hanging shoulder
15x20mm rib
Ø1957

## Slide 17: Dewar: “Pressure Vessel”
neck
3mm thick
ID: 1256
inner height: 1595
Overall height ~4770
Weight:1598kg
main vessel
Ø1118
6mm thick
ID : 1106
inner height : 2875
Ø1450
Cool down line
15x20mm rib
Ø1957
Lambda seat
55mm thick
IDs: 1250 & 1100
Ø1278x220
Centering pin
Dished head
Ø1120
8mm thick
ID : 1104

## Slide 18: Dewar: neck details
Vacuum safety valve
stiffener
6mm thick
20mm angle
Mechanical shield
thermal stiffener
6mm thick
20x60 angle
control valves
Filling line
Supercritical Ghe circuit
Collar shield
3mm thick 6082 alloy
LN2 circuit
Centering port
(shipment)
thermowells

## Slide 19
Flange: 756kg
Thickness: 65mm
Flange: 406kg
Thickness: 23mm
Height: 168
Lambda Insert: top flange new design for hydraulic test 7.15 bara

## Slide 20: Lambda Insert
Ø1450
Lhe subCooler Hx681
Titanium TieBeam Ø42,4x3,2 H1563
Mechanical shield
Level probe
Level probe
SV683 300mbar
pumping
FV680/SV681 100mbar
filling
SV680 3bar
Thermal shields
Vacuum Box (HX680+JT valve)
Open bath outgassing

## Slide 21
Pumping port
130 pipes 14/16
Pumping port
Filling port
H:820
pumping
filling
Ø354
Height 1283
Level port
Inner surface ~1m²
43kg
HX680
Filling valve
JT valve
Radius: 400&535
H340mm
Vacuum Box - HX681

## Slide 22: Double Reheater - Hnoss Reheater
1 low pressure circuit Ø114
2 atm pressure circuits : Ø40
Ø800
H2993mm
Weight 670kg
2 low pressure circuits Ø114
3 atm pressure circuits : Ø60
Ø1250
H2993mm
Weight 1140kg
Lines adaptation to study !!