# RF tests of the dressed cavity

Vitaliy Goryashko, Han Li

The list of tests in some order of priority reads

- Central cavity frequency and spectrum of higher-order modes (HOM) +

- Loaded Q (basically measurement of the 3 dB bandwidth)  +

- Q0 (calorimetric measurement) +

- Max gradient  +

- Tuning range of the slow step tuner +

- Stabilization of the cavity field with LLRF using only RF compensation - +

- Dynamic Lorentz force detuning (change in the cavity frequency during the RF pulse) -

- Compensation for the dynamic Lorentz force detuning with the fast piezo tuner -

- Stabilization of the cavity field with LLRF using both RF and piezo tuner compensation -

- Onset and level of field emission +

- Sensitivity to helium pressure fluctuations +

- Multipucking  +

- Overall test of electronics. - +

All these tests are applied to the cold cavity but when the cavity is warm, it is also important to transmit an RF signal through the cavity to ensure that RF connections are functional.

The tests 1 and 2 are performed with a VNA.

The tests 3 and 4 can conveniently be done with a pulsed self-excited loop (SEL) because it naturally tracks the cavity frequency and keeps the cavity gradient at a constant level set by the setting of SEL.

The test 5 requires a separate electronics provided by IPNO.

In the test 6, the cavity has to be driven in the so-called generator-driven configuration, which implies an RF feed at a constant frequency. The excited in the cavity field must be stabilized with the Lund LLRF.

The tests 7, 8 and 9 also require special functionality of the Lund LLRF, which is not available at the moment. The test 7 can in principle be done with the FREIA LLRF.

The tests 10-12 again can conveniently be done with a pulsed SEL. In particular, in the test 11 we track the cavity frequency with SEL and in parallel measure fluctuations of helium pressure.

The color coding green, yellow and read means the degree of readiness for corresponding tests.