## Slide 1: Test plan of ESS HB elliptical cavity
Han Li
On behalf of FREIA team
FREIA Laboratory, Uppsala University
22th of Sep. 2017

## Slide 2
Central cavity frequency and spectrum of higher-order modes (HOM) +
Loaded Q (basically measurement of the 3 dB bandwidth) +
Q0 (calorimetric measurement) +
Max gradient +
Tuning range of the slow step tuner +
Stabilization of the cavity field with LLRF using only RF compensation - +
Dynamic Lorentz force detuning (change in the cavity frequency during the RF pulse) - +
Compensation for the dynamic Lorentz force detuning with the fast piezo tuner -
Stabilization of the cavity field with LLRF using both RF and piezo tuner compensation -
Onset and level of field emission +
Sensitivity to helium pressure fluctuations +
Multipacting +
Cryo related test both at 4 K and 2 K +
Overall test of electronics. - +
The list of tests in some order of priority

## Slide 3
VNA
SGD
signal
generator driven
SEL
Lund system
Lund university
CRYO
Warm test | Cool down | Cold test | Warm up
Central cavity frequency and spectrum of HOM Qe | Frequency shift due to cool down | Coupler cold conditioning | Frequency shift vs. T
 |  | Cavity conditioning | 
 |  | Central frequency and HOM Loaded Q and Qe | 
Coupler warm conditioning |  |  | 
 |  | Cavity level profile: let the LHe evaporate to low levels Effect of CV105 in heat load Cavity's power limit Effect of different FPC cooling temperatures in heat load Max load on the 2K pumps | 
 |  | Q0 Dynamic heat load Max gradient Dynamic Lorentz force detuning | 
 |  | Stabilization of the cavity field with LLRF using only RF compensation Dynamic Lorentz force detuning Tuning range of the slow step tuner | 
 |  | Tuner related testing | 

## Slide 4
Warm test detail |  |  |  | 
Process or check item | Work Description | Team | Loop and special device | Power
Build loop and cabling | Build SEL | RF |  | Low power
instrumentation check | validation of instruments SEL, FPGA software debuging | RF | SEL: Copper cavity, SEL, FPGA, interlock switch, RF switch and RF station | 
 |  |  | SGD: Signal generater driven system, interlock system, coupler vacuum, RF station, conditioning software | 
 |  |  | VNA: VNA, frequency auto checking software | 
S parameter measurement | Get central cavity frequency | RF | VNA | 
Qe | Simply estimate the coupling factor and Qe | RF | VNA | 

## Slide 5
Warm test detail |  |  |  | 
Process or check item | Work Description | Team | Loop and special device | Power
preparing of Interlock system | Validation of software Set related parameters in the software and vacuum interlock thredshold | CTRL RF | SGD: software, RF station, | Low power
Coupler warm conditioning | Ramp up forward power to 1.1 MW? Increase pulse width to 3.5ms ? Get an optimal procedure for RF conditioning | CTRL RF | SGD: RF station, FREIA software | High power

## Slide 6
Cool down test detail |  |  | 
Process or check item | Work Description | Team | Loop and special device
Frequency shift vs. T | Frequency shift due to cool down | RF | VNA, frequency auto checking software
Cool down to 4 K |  | CRYO | 
Cavity level profile | let the LHe evaporate to low levels (below 30%) | CRYO | 
Frequency shift vs. Pressure | Frequency shift due to pressure | RF | VNA, frequency auto checking software
Cool down to 2 K |  | CRYO | 
Static heat loads | Effect of CV105 in heat load Cavity's power limit: apply heat power until the level and/or the pressure in the cavity shows instabilities Effect of different FPC cooling temperatures in heat load | CRYO | 
Max. 2K pumps capacity at 2K | Max. 2K pumps capacity at 2K (theoretically at 90W) | CRYO | 

## Slide 7
Cold test detail (I) |  |  |  | 
Process or check item | Work Description | Team | Loop and special device | Power
Coupler Cold conditioning | Conditioning coupler with detuned cavity or with a frequency outside the cavity bandwidth Ramp up forward power to 1.1 MW? Increase pulse width to 3.5ms? | RF CTRL | SGD: RF station, FREIA software Data acquisition system Personal safty system | High power
Dressed cavity packet cold conditioning | Conditioning coupler and cavity at resonant frequency Ramp up forward power to 232KW? Increase pulse width to 3.5ms? | RF CTRL | SEL: RF station FPGA, Interlock system, RF switch Data acquisition system Personal safty system | 
Central frequency and HOM Loaded Q and Qe | Get central cavity frequency and high order mode Loaded Q and -3dB bandwidth Qe | RF | VNA | Low power

## Slide 8
Cold test detail (II) |  |  |  | 
Process or check item | Work Description | Team | Loop and special device | Power
Q0 Dynamic heat load Max gradient Dynamic Lorentz force detuning | Runing pulse SEL from low power with 3.5ms, 14Hz forward pulse Get average Q0 value as a function of average gradient Checking the radiation Increas the forward power until reaching the nominal gradient Measure dynamic heat load at different power level Use I/Q meadth to check the frequency shit during the pulse | RF CRYO | SEL: SEL, FPGA, RF station Interlock system, RF switch, Helium gas flowmeter, Data acquisition system, Personal safty system | High Power
Dynamic Lorentz force detuning | Use Lund system to generator step pulse input ? Use FPGA to check the frequency shit during the pulse | RF | Lund system , FPGA, RF station Data acquisition system, Personal safty system | 

## Slide 9
Cold test detail (III) |  |  |  | 
Process or check item | Work Description | Team | Loop and special device | Power
Stabilization of the cavity field with LLRF using only RF compensation | Use Lund system to produce input pulse Ramp up power level from low power to 232KW Stay at nomial gradient for a certain time to confirm the stabilization of the cavity field | CRYO RF | Lund system RF station Data acquisition system, Personal safty system | High power
Tuning range of the slow step tuner | Slow tuning control by lund system Slow tuning control by IPNO software ? | CRYO RF | Lund system Data acquisition system, | Low power
Tuner related testing | Test that lund university are interested ? | CTRL Lund | Lund university | 

## Slide 10
Warm up test detail |  |  | 
Process or check item | Work Description | Team | Loop and special device
Frequency shift vs. Pressure | Frequency shift due to pressure | RF | VNA
Warm up to 4 K |  | CRYO | 
Frequency shift vs. T | Frequency shift due to warm up | RF | VNA
Warm up to 300 K |  | CRYO | 

## Slide 11
Time schedule |  |  |  | 
Process or check item | Team | Loop and special device | Power | Time
Build loop and cabling | RF | change componets which adaopts to 704 MHz | Low power | 1 week
instrumentation check | RF | SEL: Softwer, FPGA, interlock switch, RF switch and RF station |  | 2-3 weeks
 |  | SGD: Signal generater driven system, interlock system, coupler vacuum, RF station, conditioning software |  | 
 |  | VNA: VNA, frequency auto checking software |  | 
S parameter measurement | RF | VNA |  | 0,5 day
Qe | RF | VNA |  | 0,5 day
Time schedule

## Slide 12
Time schedule |  |  | 
Process or check item | Team | Loop and special device | Time
Frequency shift vs. T | RF | VNA, frequency auto checking software | 1 week
Nitrogen cool down | CRYO |  | 1day
Cool down to 4 K | CRYO |  | 0,5 day
Static heat loads | CRYO |  | During weekend
Cavity level profile let the LHe evaporate to low levels (below 30%) | CRYO |  | 0,5 day
Frequency shift vs. Pressure | RF | VNA, frequency auto checking software | 1 days
Cool down to 2 K | CRYO |  | 1 days

## Slide 13
Time schedule |  |  |  | 
Process or check item | Team | Loop and special device | Power | time
preparing of Interlock system | CTRL RF | SGD: Conditioning software, RF station | Low power | 2 days
Coupler Cold conditioning | RF CTRL | SGD: RF station, FREIA conditioning software, Data acquisition system, Personal safty system | High power | 1 week
Dressed cavity packet cold conditioning | RF CTRL | SEL: RF station, FPGA, interlock system, RF switch, Data acquisition system, Personal safty system |  | 
 |  |  |  | 1 week
Central frequency and HOM Loaded Q and Qe | RF | VNA | Low power | 0,5 day

## Slide 14
Time schedule |  |  |  | 
Process or check item | Team | Loop and special device | Power | time
Effect of CV105 in heat load Cavity's power limit Effect of different FPC cooling temperatures in heat load | CRYO | SEL: SEL, FPGA, RF station Interlock system, RF switch, Helium gas flowmeter, Data acquisition system, Personal safty system | High Power | 
Q0 Dynamic heat load Max gradient Dynamic Lorentz force detuning (with and without tuner contacted) | RF CRYO |  |  | 2 weeks
Dynamic Lorentz force detuning (with and without tuner contacted) | RF | Lund system , FPGA, RF station Data acquisition system, Personal safty system |  | 1 day

## Slide 15
Time schedule |  |  |  | 
Process or check item | Team | Loop and special device | Power | time
Stabilization of the cavity field with LLRF using only RF compensation | CRYO RF | Lund system RF station Data acquisition system, Personal safty system | High power | ?
Use Lund system to produce input pulse Ramp up power level from low power to 120KW Stay at nomial gradient for a certain time to confirm the stabilization of the cavity field |  |  |  | 
Tuning range of the slow step tuner | CRYO RF | Lund system Data acquisition system, | Low power | 2 days?
Tuner related testing | CTRL Lund | Lund university |  | 

## Slide 16
16
Time schedule |  |  | 
Process or check item | Team | Loop and special device | time
Frequency shift vs. Pressure | RF | VNA | ?
Warm up to 4 K | CRYO |  | 1 hour
Frequency shift vs. T | RF | VNA | ?
Warm up to 300 K | CRYO |  | 5 days