## Slide 1: Quench Protection
Younguk Sohn
FREIA, Meeting for H-β Elliptical cavity test, October 3, 2017

## Slide 2: Beam Dump by Quench 2013/12/01/19:17
Reflected power (Pref) decreased, forward power (Pfor) increased slightly, but accelerating voltage (Vacc) decreased. Then it was typical quench !!
P_tran: Vacc
P_for
Beam current
P_ref

## Slide 3: Stored Energy in 500 MHz Cavity , PLS-II

## Slide 4: Beam Dump 2013/12/01/19:17
LHe evaporation rate
Pressure @He vessel
Cavity vacuum pressure
Quench is happened, then
→ All RF power is reflected to load and transferred to beam
→ Only stored energy in cavity is dumped to normal zone
(~ ten Jules)
→ Cavity temperature increases
→ LHe evaporation rate increases
And cavity vacuum become bad
Beam current

## Slide 5: Beam Dump 2013/12/01/19:17
Temperature @cavity bottom
Temperature @cavity top

## Slide 6: Algorism of Quench Protection
LLRF
Run mode (beam): 1
Processing mode: 0
Forward power limit
Processing mode :
Vreflected - freflected (%) * Vforward < 0 (if P_reflected decreases)
where freflected (%) is set value, the higher the faster response.
Run mode (Beam operation):
Vprove - fprove (%) * Vset < 0 (If Vacc decreases)
Vreflected - freflected (%) * Vforward < 0
Forward power limit
- Vforward > Vlimit
where Vlimit is set value
Quench trigger time is dependent on quench intensity, fastest time is < 3μs
※ V??? is direct signal from power meter