## Slide 1: Teleconference ACS – FREIA 31st October 2017
Agenda:
Manufacturing follow-up
Control and command follow-up
Planning
Commissioning test organization

## Slide 2: Manufacturing follow-up
Cryostat
  Dewar piping completed with insulation (W43)
  Thermal shield : cooling pipes validated (He leak test and thermal shock), welding ongoing (W?)
  Instrumentation and wiring ongoing (started W43)
    Technergie visit planned Tuesday to install the junction box (W44)
Magnet insert
  Manufacturing drawings reviewed (26th October)
  Top lid (starting part) is under plasma cutting
    Issue to keep the tolerances after delivery to the subcontractor so send for water jet cutting instead
  Current lead equipment (flowmeters, heaters, sensors) to be defined by ACS
    Need date on the pressure drop for the current lead cooling
  Heat exchanger HX683 welding design presented and validated by CD and ACS
  Manufacturing start in W44

## Slide 3: Manufacturing follow-up
Liquid insert
  Waiting for final cabling
  Insertion inside the cryostat in W44
Valve box
  Piping done, test ongoing to find and repair any leak before closure (W44)
    Leak found on the top thermal screen repaired but not confirmed
  Heaters and thermometers set on the thermal shield, 4k pot and heat exchanger
  Cernox installation ongoing

## Slide 4: Manufacturing follow-up
Transfer lines
  Manufacturing of the single line and multiline under way
Simulator box
  Manufacturing drawings reviewed, translation to be checked before sending
  Manufacturing started W43

## Slide 5: Control and Command follow-up
Control system design
  Programming done, waiting for test with GERSEMI systems
Control system procurement
  Junction boxes done
    Missing connectors BURNDY male 19 pins, to be send by UU
  Electrical cabinet delivered at CD
Cabling of the VB to the cabinet in W44
Cabling of the VC to the cabinet in W45
  Cables cut after installation on dummy trays, reproducing the laboratory FREIA infrastructure

## Slide 6: Planning: GERSEMI
 | 2017 |  |  |  |  |  |  |  |  |  |  |  |  | 2018 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
 | October |  |  |  | November |  |  |  |  | December |  |  |  | January |  |  |  | February |  |  |  | March |  |  |  |  | April |  |  | 
 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47 | 48 | 49 | 50 | 51 | 52 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17
junction box manufacturing |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
VB electrical wiring, closure and test |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Liquid insert electrical wiring |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
VC electrical wiring and test |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Simulator manufacturing and test |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Simulator electrical wiring |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Magnet insert manufacturing |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Magnet insert wiring |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Electrical cabinet delivery |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Final electrical wiring of control cabinet & test |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Control cabinet & automaton test |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
- Factory acceptance test with VC |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
- Factory acceptance test with simulator |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Control cabinet & automaton test with magnet insert |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Shipping of the vertical cryostat and liquid insert |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Shipping of the valve box and simulator |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Installation of project onsite |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Cryogenic tests with simulator |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Final installation of project without simulator |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Acceptance tests with liquid insert |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Acceptance tests with magnet insert |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 

## Slide 7: Commissioning test organization
Preliminary tests
  Final assembly tests:
    Leak tests pipes and valves
Commissioning tests
  Week 47-48: wiring check VC and liquid insert
  Week 49: sequence tests with automaton and VC (22,23,25)
  Week 50-51: sequence tests with automaton and simulator (1,3,5,8,9,10,12,14,21,23)