## Slide 1: FREIA Laboratory Facility for Research Instrumentation and Accelerator Development
Accelerator Development
in Uppsala
Roger Ruber
Uppsala, 11 June 2019

## Slide 2: Uppsala Accelerator History
1477: Uppsala University, oldest in Scandinavia
25'000 students, 7'000 staff
historical profiles: Linné, Rudbeck, Celsius, Ångström, Svedberg
1940's: The(odore) Svedberg builds a cyclotron
Gustaf Werner synchro-cyclotron (1947 - 2016)
  nuclear physics & oncology
CELSIUS ring (1984 - 2005)
  nuclear & particle physics
2000's: External projects
CTF3/CLIC (since 2005)
FLASH/XFEL (since 2006)
ESS (since 2009)
2010's: New ventures
FREIA laboratory (est. 2011)
Skandion clinic (est. 2015)

## Slide 3: FREIA Laboratory
cryogenics
- liquid helium
- liquid nitrogen
control room
- equipment controls
- data acquisition
radio-frequency (RF) power sources
3 bunkers
with test stands
horizontal cryostat
vertical cryostat
State-of-the-art Equipment
Facility for Research Instrumentation and Accelerator Development
Funded by KAWS, Government, Uppsala Univ.
Competent and motivated staff
collaboration of physics (IFA) and engineering (Teknikum).

## Slide 4: Accelerator Research & Development
High Intensity Proton Beams
Ultra Bright Electron Beams
SC Cavities & Magnets
RF Generation & Control
Cryogenics & Test Stands
Accelerator
Physics
Accelerator
Technology

## Slide 5
Overview of Cryogenic Test Stands
cryogenic test stands

## Slide 6: Cryogenics
Helium liquefaction
  150 l/h at 4.5K (LN2 pre-cooling)
  2000 l LHe dewar/buffer, 3+1 outlets
  cryostats connected in closed loop
Gas recovery
  100 m3 gasbag
  3x 25 m3/h compressor
  10 m3 200 bar storage
2K Pumping
  ~3.2 g/s at 10 mbar
  ~4.3 g/s at 15 mbar
  110(90)W at 2.0(1.8)K
Liquid nitrogen
  20 m3 LN2 tank
Sub-atmospheric pumping station
Helium gas recovery system

## Slide 7: “HNOSS” Horizontal Cryostat
HNOSS = Horizontal Nugget for Operation of Superconducting Systems
Test of superconducting cavities/devices
  3240 x ø1200mm inner volume
  up to two cavities simultaneously,
  each equipped with helium tank,
Low or High power RF testing
  fundamental power coupler (top, bottom, side)
  (cold) tuning system
Operation in the range 1.8 to 4.5K.
Available as ARIES
Trans National Access Facility

## Slide 8: “Gersemi” Vertical Cryostat
Under commissioning
Test of SC cavities & magnets (<350kJ)
  3.2m x ø1.1m total volume
  2.65m x ø1.1m below lambda plate
    design includes joint for lambda plate
Three operation modes
  vacuum; liquid bath; pressurized (bath with 2K heat exchanger)
Operation in the range 1.8 to 4.5K
Available as ARIES
Trans National Access Facility
2 x 2000 A
Presentation by K. Pepitone
"News from the FREIA test facility"
Presentation by R. Santiago Kern
"Challenges of a 1.9 K vertical cryostat design"

## Slide 9
Development of High Power RF Technology
RF Generation & Control

## Slide 10: High Power RF Amplifiers
400 kW pulsed (352 MHz)
  2 stations, each 2 tetrodes TH595(A)
  3.5 ms, 14-28 Hz
  ESS prototype development
50 kW CW (352/400 MHz)
  single tetrode TH571b
1.2 MW pulsed (704 MHz)
  HV modulator & klystron test for ESS
CERN 50 kW CW

## Slide 11: Solid State RF Amplifier
Transistor Amplifier Module
single ended RF power amplifier
based on BLF188XR
1250 W and 70% efficiency
Amplifier Demonstrator
8 modules, 10.5 kW
69% efficiency
  pulsed 14 Hz, 3.5 ms (ESS)
L. Haapala et al. Electronics Letters, vol. 52, (2016) p.1552.
D. Dancila et al. IOP Conf. Series, J. Physics Conf. Series vol. 874, 2017 (012093).

## Slide 12: Low-loss Compact Combiners
Compact Cavity Combiner
352 MHz 200 kW
  12 input ports
  0.2% insertion loss
Compact Planar Combiner
352 MHz 10 kW
  8 input ports
  Gysel type
  line coupling compensates parasitic coupling
352 MHz 20 kW
  2 to 1
    2 ext. loads
  combiner/splitter
  insert.loss 0.1 dB
V. A. Goryashko et al. IEEE Microwave and wireless components Letts, vol. 28, 2018.
M. Jobs et al. IEEE Trans. Components, Packaging Manufacturing Tech., vol. 8, 2018.
L. Hoang Duc et al. J. of Engineering, 2017.
Input RF power
Output RF power

## Slide 13: FPGA Based Digital LLRF
Signal Driven
2 ADC inputs at 250 Msps
  (*) analogue bandwidth of 750 MHz
2 DAC outputs at 500 Msps
Digital downconversion to baseband 0 Hz, no analog mixers
  downconverted signal at 10 Msps or 1 Msps, selectable
undersampling to operate at any frequency from 10 to 750 MHz*
Self-excited Loop
CW or
pulsed mode
  switch closes the loop for a duration of 2.86 ms, repetition rate of 14 Hz.
Code in the FPGA

## Slide 14
Development of SRF cavities for ESS
Superconducting cavities

## Slide 15: ESS Cavities & Cryomodules
Double Spoke Cavity, 352 MHz
Prototype cavity
  without and with FPC
  RF conditioning
  Q0, gradient, fill time,
  Lorentz force detuning, microphonics
  test LLRF, SEL,
  tuner operation
  nominal gradient
Cryomodules
  prototype valve box & cryomodule
  13 series cryomodules
    Oct. 2019 – end 2020 (~6 weeks/CM)
Elliptical Cavity, 704 MHz
RF stations
  acceptance test of HV modulator for ESS local test stand
  test RF distribution (circulator, load)
Prototype high beta elliptical
  with power coupler and tuner
  RF conditioning
  Q0, gradient, fill time, heat load
  Lorentz force detuning, microphonics
  test LLRF, SEL, tuner operation

## Slide 16: Spoke Cryomodule Prototype (2019)
Warm RF conditioning
  ~3 days/cavity
  MP bands were consistent with HNOSS test
    strength depends on pulse length,
    1st/2nd conditioning…
Cold RF conditioning
  no coupler activity
  Quench during cavity conditioning at 4 K
    burst disc rupture → thermal cycling
Cavity #2 performance
  multipacting regions similar as prototype
    2-3; 4-5; 7-8 MV/m
  field emission sensitive to tuner motion or position (under investigation)
FREIA CAV2 after tuner test (2019)

## Slide 17: Spoke Cryomodule Prototype (2019)
Fast tuner performance
  Lorenz force detuning compensation (piezo)
w/o correction
w/ correction
fshift
Pfwd
Pref
Ptrans

## Slide 18
Test and Development of SC Magnets for CERN
superconducting magnets

## Slide 19: Orbit Correctors for HL-LHC
Test of nested dipole orbit corrector magnets for the High Luminosity upgrade of LHC
  magnet design and construction by CIEMAT (Spain)
  test at FREIA (20 magnets)
Parameter | Unit | Value
Aperture | mm | 150
Magnetic length | m | 1.2
Nominal integrated field | T m | 2.5
Nominal current inner | A | 1625
Nominal current outer | A | 1474
Short sample current at 1.9 K | A | 
Loadline fraction at 1.9 K |  | 
Strand diameter | mm | 0.480
Cu/no_Cu |  | 1.75
Number of turns per layer inner |  | 140
Number of turns per layer outer |  | 191
Nominal strand current density | A/mm2 | 
Nominal superconductor current density | A/mm2 | 
Nominal differential inductance inner | mH | 58.4
Nominal differential inductance outer | mH | 118.8
Nominal stored energy inner | kJ | 87
Nominal stored energy outer | kJ | 150
Maximum hotspot temp | K | 270
MCBXF

## Slide 20: Magnet Powering & Quench Monitoring
2 x 2000 A
Data acquisition
Computer
PLC Slot
Quench detection system
Patch panel
Data acquisition
slot
Computer
IGBTs based EE system
Presentation by K. Pepitone
"Magnet safety system and supervision in FREIA"

## Slide 21: CCT Magnet Development
Canted-Cosine-Theta magnet is a dipole based on the superposition of two oppositely skewed solenoids with respect to the bore axis.
  produces a perfect cosθ field,
  is cost effective compared to a conventional SC dipole
    but not the same field strength possibilities
CERN (G. Kirby)
Scanditronix (A. Ahl)

## Slide 22
CLIC, Free Electron Lasers, Ultra-fast Electron Diffraction
high brilliance electron beams
CLIC
FELs
Ultrafast Electron Diffraction

## Slide 23
European Spallation Source (ESS) and High Luminosity LHC (CERN)
high brilliance proton beams
ESS Superconducting Linac
HiLumi LHC
ESS Neutrino Super Beam

## Slide 24: Summary
Uppsala University & FREIA Laboratory actively developing accelerator and instrumentation technology
Technology Development
NC and SC RF cavities
SC magnets
RF power generation
LLRF and controls
Physics Research
high brilliance beams
superconducting RF
RF breakdown
Academic Teaching