## Slide 1
Marcos Turqueti
US Magnet Development Program
Lawrence Berkeley National Laboratory
3nd International Magnet Test Stand Workshop
Magnet Test Facility at LBNL

## Slide 2
Outline
Maxim Marchevsky – Physicist
Marcos Turqueti – Electrical Engineer
Jordan Taylor – Electrical Technician
Jim Swanson – Cryo Technician
Daniel Davis – Ph.D Student
MTF team:
Facility Overview
CLIQ
Extraction
Power Supply
CryoDAQ

## Slide 3
The Facility
Magnet Test Facilities
It is located in the Berkeley Hills near Berkeley, California;

## Slide 4
The Facility
The Magnet Test Facility (MTF) at LBNL has the mission of testing and characterizing prototype high field superconducting magnets;
Several magnet parameters, such as current, training behavior, and field quality are tested and characterized at MTF;
The facility is capable of adapting to an wide variety of magnets and provide unique instrumentation capabilities;
Several upgrades are current ongoing at LBNL to further improve the Magnet Test Facility capabilities.
Current scheduled for testing (July):
  Gantry, CCT5, C2B

## Slide 5
Upgrades - Coupling-Loss Induced Quench system (CLIQ)
Max: 9.9 kJ, 800 V, 45 mF

## Slide 6
6
Bleed resistor for each capacitor.
Normally closed relay
IGBT gate control
Capacitor bank
voltage measurement
Capacitor bank charging
IGBT switching on/off capacitor banks.
Upgrades - CLIQ

## Slide 7
Energy Extraction System
Silicon Controlled Rectifier (SCR) based extraction system. Four SCRs capable of switching 6kA (55 ºC) each.
  Maximum tested current rise 500A/s
  Maximum voltage of 1350V
  Maximum SCR operating temperature of 125C, with interlocks set to 90C
  Turn off time of 0.5ms
9kJ capacitor bank necessary to disengage SCR’s during energy extraction with energy proportional to maximum extraction current.

## Slide 8
Current transducer LF 2010-S, IPN=2 kA
IGBT, 2.4 kA, 1.7 kV
IGBT assembly with cooling
Dump resistor unit – specifications next page.
By J. Taylor
Bus bars
Dripping pan
Four Soresen (10 V, 1.2 kA)
Upgrades - Energy Extraction
Scenerio | Resistance (ohm) | Energy rating (kJ)
1 | 0.1 | 400
2 | 0.05 | 500
3 | 0.0333 | 600
4 | 0.025 | 700
5 | 0.02 | 800
IGBT Based System

## Slide 9
Main Power Supply
New 25kA power supply (Transferred; Los Alamos)
Maximum current: 25 kA (limited to 23.5 kA by voltage drop on the lines + SCRs balance resistors)
Maximum voltage: 20 V
Remotely controlled from the MTF control room via Ethernet interface (dedicated network line)
SGI Series DC
5kA secondary magnet power supply is now in operation

## Slide 10
Quench Detection System
Controller
Chassis
VxWorks
LabVIEW RT
LabWindows/CVI RT
LabVIEW FPGA
Custom VHDL
NI-RIO
Win/Linux
LabVIEW
LabWindows/CVI
NI-9215
4 ch 100kS/s 16bit
NI-9401
8ch 100ns TTL
Modules
PowerPC
FPGA
LBNL QUENCH DETECTION SYSTEM

## Slide 11
Quench Detection System
Ww – Window width 2us- 100ms
TH1 – Trigger Threshold
TH2 – Action Threshold
SIGNAL ORIGIN

## Slide 12
Active Cryogenic Electronic Envelope (2013)
  U.S. Patent No. 10,240,875 (Supercon)
Commercial off-the-shelf cryogenic components (4k-77k)
ADCs, Amplifiers, Diodes, FPGAs, Laser Diodes, Optical Isolators , Voltage Regulators, Transistors and Passives
Field Programmable Gate Array (FPGA) on Cryogenic Environments
FPGA operating with COTS ADC
FPGA based ADC
Cryogenic Magnetic Sensors
  Acoustic Sensors (Maxim)
  Electron Beam Sensor (LDRD)
160nm Cryogenic ASIC
16 channels 12b 10MSPS ADC, 1 MSPS 10b ADC, Charge Amplifier, Voltage Regulator, Strain Gauge, Sea-of-Transistors, Memory (C.Grace, D.Gnani)
New Activities : CryoDAQ R&D

## Slide 13
Strain Gauge
Voltage Taps
Temperature Sensor
New Activities : CryoDAQ R&D

## Slide 14
Architectures for cryogenic DAQ
FPGA
ADC
AMP
MUX
Sensors
OPTO
FPGA
ADC
AMP
ADC
ADC
Power Reg
Power Reg
TX/RX line
TX/RX line
Power:
Switched Capacitor+LDO
Power over Fiber
LED supply
Communications:
JESD
LVDS
CMOS
Ethernet (PHY?)
Over copper or fiber
Single or multiple lanes.
Typically one lane provides 300Mbps
Isolation:
Opto or capacitive
Sensors

## Slide 15
ASICs for Cryogenics Design by LBNL

## Slide 16
Systems with COTS cryogenic components
LN Test
LHe Test
Slow Amplifier
(Acoustic Instrumentation)
Fast Amplifier -500MHz
Basic COTS component
ADCs
  24b / 20kSPS
  12b / 10MSPS
  4b / 500MSPS
Amplifiers (up to 500MHz)
Diodes (HV protection 1A)
FPGAs (28nm- 45nm)
Laser Diodes Optical Isolators
Voltage Regulators
  LDO and SW capacitor
Transistors
  JFET
  HEMT
  MOSFET
Passives
  Capacitors and resistors

## Slide 17
Reconfigurable Flash Folding Architecture. With this architecture it is possible to change the ADC resolution on the fly.
FPGA Based n-bit Analog to Digital Converter