## Slide 1: News from the GSI test Facility @ CERN
Antonella Chiuchiolo on behalf of GSI-CERN collaboration on Super-FRS testing
3rd International Workshop of the Superconducting Magnets Test Stands,
Uppsala 11-12 June, 2019

## Slide 2: GSI accelerators complex
Super-FRS
SIS100
SIS18

## Slide 3: Outline
Super-FRS superconducting magnets
GSI Test Stand layout @ CERN
Facility infrastructures
Facility status today

## Slide 4: Super-FRS superconducting magnets
2019-06-11
# | Type | Mass | Cold mass | LHe volume | Operation current | Stored energy | Inductance | Aperture | Peak field
 |  | (kg) | (kg) | (m3) | (A) | (MJ) | (H) | (mm) | (T)
21 | Dipoles | 50000 | 2000 | 0.025 | 236 | 0.45 | 26 | 70 x 190 | 1.6
7 | Short Multiplet | 28000 | 20000 | 0.8 | 300 | 1 | 21 (Quad) | 380 | 2.5
24 | Long Multiplet | 60000 | 45000 | 1.5 | 300 | 2.7 | 21 (Quad) | 380 | 2.5
4 m
2.7 m
2.5 m
Standard Dipoles
Short Multiplets
From 2 up to 9 magnets in the same cold mass
7 m
Branch dipoles
Long Multiplets

## Slide 5: Super-FRS magnets test stand @ CERN
CERN’s building 180
Preparation area
Test area
Power converters
Control racks
Cold Box
Dewar
Control racks
Facility Construction
2012-2018  CERN contribution
180
2019-06-11

## Slide 6: Super-FRS magnets test stand @ CERN
CERN’s building 180
Preparation area
Test area
Power converters
Control racks
Cold Box
Dewar
Control racks
INFRASTRUCTURE
3 test benches
3 cryogenic distribution systems
9 power converters circuits
3 energy extraction systems
2 over-head cranes (40T and 60T)
OPERATION
Imax = 330 A
Top = 4.5 K
Facility Construction
2012-2018  CERN contribution
180
2019-06-11

## Slide 7: Testing team and plan
Facility Commissioning: 2018-2019 with the first magnet
Facility Operation: 2018-2023 GSI team + CERN Cryo OP
2019-06-11

## Slide 8: Test facility interfaces
Interface | GSI | CERN
Electric | Current lead warm terminal | Cable from power converter
Mechanic | Magnet feet | Spacer on the floor
Cooling | Cryogenic jumper line | Satellite valve box
(Helium return) | Instrumentation panel | Helium return line
Vacuum | Cryostat vacuum pump port | Vacuum pump
Alignment | Fiducial targets on cryostat | Fiducialization
Instrumentation | Cabinet for sensor signals | Control system
Magnetic | Magnet | Magnetic measurement systems
Facility
2019-06-11

## Slide 9: Cryogenics
Cryogenic main sub-systems
  Pre-cooling system
  LHe system
  Cryogenic distribution system
Courtesy of G. Rolando (TE-CRG-ME)
150 g/s at 18 bar
50 g/s at 17 bar
2019-06-11

## Slide 10: Power Converter, Energy Extraction, Power Cables
3 x [±500A/±120V] for the main circuits + EE system with 2.8 Ohm resistance for the quadrupoles
6 x [±600A/±40V] for the corrector circuits
9 load switches interlocked
78 DC powering cables, with a combined total length of about 2900 m
Power converters Racks
Load Switches
Power Cables
2019-06-11

## Slide 11: Universal Quench Detection System
Front-end channels
(up to 16)
Galvanic insulation
2.5kV
Redundant power supplies
(monitored)
Digital mid-plane
Interlocks & communication
Field Bus
Backplane with signal connector(s)
Redundancy on the level of crate
    Enhanced availability of supervision
Qualified in radiation
    Radiation tolerant up to 100 Gy
Courtesy of M. Zerlauth (TE-MPE-MS)
base line prototype for the HL-LHC quench detection system
Integrated data acquisition system oﬀers high sampling rates
2019-06-11

## Slide 12: Interlock PLC
Power commutators
Power converters
EE switches
QDS
Platform access
Cryogenic PLC
2019-06-11

## Slide 13: Magnetic measurements
Magnetic field measurements for magnet acceptance
4 measurements devices
Giancarlo Golluccio’s talk: Magnetic measurements techniques for Super FRS Today at 17:55
2019-06-11

## Slide 14: GSI contribution
Satellite valve box
(CERN)
Jumpler line
(GSI)
Electrical Cabinet
Instrumentation panel
2019-06-11

## Slide 15: GSI contribution
SCADA software based on the UNICOS-framework: monitoring and control of magnet and current leads instrumentation
Configuration of the different panel types for each magnet
2019-06-11

## Slide 16: Status Today
First SM arrived on Feb 7th, 2019
Electrical integrity Tests done
SM preparation for the bench done
Installation on the test bench May 7th, 2019
Cold mass survey measurements
Insulation vacuum and leak test on going
1st Cooldown 25th June, 2019
2019-06-11

## Slide 17
Thank you for your attention

## Slide 18
Back up

## Slide 19: Test plan for FoS short multiplet
Steps
magnet welcome
magnet on lorry
magnet at preparation area
magnet at test bench
isolation vacuum pumping
cool down
cold test without excitation
cold test wit excitation
  Magnetic measurements
warm up
warm test under vacuum
warm test at atmosphere
magnet removal preparation
magnet at preparation area, transport preparation
magnet goodbye
2 times
(thermal cycle)
(Delivery at CERN, 7th Feb. 2019)
(Shipment to FAIR, end of 2019)
Now