## Slide 1: Mechanical measurements in test stands Optical fibers and Cloud technology M. Guinchard
12th June 2019

## Slide 2: Outline
Introduction
  Optical strain measurements based on Bragg® techniques
    Validation process
    Magnet experience
  Digital transformation

## Slide 3: Introduction
Why perform mechanical measurements ?
Why mechanical measurements in superconducting magnets ?
  Support the mechanical design during prototyping phases;
  Validate Finite Element Analysis (FEA) during assembly phases, thermal cycles (down to 1.9K), powering tests and warmup;
  Check the integrity of the structure during the assembly process.
Inputs for FEA :
Thermo physical properties
Mechanical properties
FEA validation & Safety

## Slide 4: Introduction
How to perform mechanical measurements ?
  Non contact video systems as Digital Image Correlation (DIC) to study collaring and mechanics, in particular coil peak stress during collaring on mock-up.
Camera#01
Camera#02
Lighting system
Useful techniques to understand the cinematic and to cross-check the standard instrumentation (electrical or optical strain gauges);
Limited to lab environmental conditions.

## Slide 5: Introduction
How to perform mechanical measurements ?
  Standard electrical strain gauges from 1.9K up to few hundred degrees
More than 30 years of experience of measurements in harsh conditions;
Measurement benchmark successfully done in 2017 with American lab’s;
Collaborative work with several institutes to standardized techniques and tools based on CERN experience.

## Slide 6: Introduction
How to perform mechanical measurements ?
  Optical fiber strain gauges from 1.9K up to RT based on Bragg gratings
Well known techniques at room temperature for Civil Engineering applications;
Promising techniques for harsh environments (electro magnetic disturbances);
Distributed sensors over the full fiber length;
5 Years of study based on KT agreement with industrial researchers.

## Slide 7: Optical strain measurements
Several techniques based on optical fibers (Bragg, Brillouin, Raman);
Bragg technique under validation for the moment :
  Lot of advantages except the strong coupling between  and T !
  Temperature will also affect the bonding process according to the CTE of the glue and the support material itself;
  Measurements in superconducting magnets request an absolute strain state measurement over the full lifetime including thermal cycles :
  Request a thermal compensation to compensate the apparent strain !
a : Photo elastic coefficient of the fiber
: Coefficient of thermal expansion
 : Refractive index

## Slide 8: Optical strain measurements
Temperature compensation based on an external compensator in free-free stress state :
  Excellent repeatability and reproducibility of the bonding process is required !
  Before field implementation, several validation tests need to be performed :
  Trueness determination over the temp. range to validate the measurements;
  Repeatability and reproducibility determination to validate the temperature compensation.
ISO5725

## Slide 9: Optical strain measurements
Validation process :
Tensile test machines: Zwick/Roell® Z400 for tests down to 4.2 K;
AISI 304L flat dog-bone specimens;
2 FBG/side, polyimide coated, 5 mm gage length, 60% reflectivity;
Commercial glues were used: Vishay® M610 & Araldite®;
Precision estimated by measuring the apparent strain at 77 K;
Trueness estimated by evaluation of Young Modulus (ASTM E111-04).

## Slide 10: Optical strain measurements
Precision results :

## Slide 11: Optical strain measurements
Trueness results:

## Slide 12: Optical strain measurements
Precision and trueness results:
More information :

## Slide 13: Optical strain measurements
Magnet application – MQXFS Magnet
SHT
SHB
SHL
SHR
RDZ
RAZ
RCZ
RBZ
CO204
CO205
CO206
CO203
Electrical + Optical strain gauges

## Slide 14: Hybrid system
MX840 for SG reading
MX1615 for SG reading
DI405 for FBG reading
Settings
Data
SM18 Test Server
Text files created after each quench
Mechlab Server
Catman files
Continuous storage

## Slide 15: Optical strain measurements
Magnet application – MQXFS Magnet

## Slide 16: Optical strain measurements
Magnet application – MQXFS Magnet

## Slide 17: Optical strain measurements
Magnet application – MQXFS Magnet
Assembly process at RT
Cool-down process down to 1.9K

## Slide 18: Optical strain measurements
Magnet application – MQXFS Magnet
Powering test at 1.9K

## Slide 19: Optical strain measurements
Magnet application – MQXFS Magnet
Powering test at 1.9K

## Slide 20: Optical strain measurements
Magnet application – MQXFS Magnet
 | Ratio SG vs FBG |  | 
Quench # | 203T | 205T | 206T
1 | 1.0182 | 0.9984 | 1.0191
6 | 1.0155 | 0.9974 | 1.0086
13 | 1.0057 | 1.0014 | 1.0039
18 | 1.0124 | 1.0008 | 1.006
23 | 1.0115 | 1.0039 | 1.0072
32 | 1.0132 | 1.008 | 1.012
 | R2 |  | 
Quench # | 203T | 205T | 206T
1 | 0.9997 | 0.9998 | 0.9998
6 | 0.9999 | 0.9999 | 0.9999
13 | 0.9999 | 0.9997 | 0.9999
18 | 0.9999 | 0.9997 | 0.9999
23 | 0.9999 | 0.9996 | 0.9999
32 | 0.9999 | 0.9996 | 0.9999
Powering test at 1.9K

## Slide 21: Optical strain measurements- conclusion
The precision of the bonding procedure of optical fiber sensors was established according to ISO 5725;
The trueness was evaluated down to 4.2 K: the results confirmed the capability of the optical strain sensors based on Bragg grating to measure down to 4.2 K with an accuracy below 5.5% according to ISO 5725;
A first superconducting magnet was successfully equipped with mirror instrumentation : electrical and optical strain sensors;
Measurements done during assembly, cooldown and powering phases confirmed the validity of the optical strain sensors for superconducting magnet applications;
  Additional studies are on going to improve the fiber protection, splicing operation for cryogenic applications, etc…

## Slide 22: Digital transformation
The Digital Transformation Revolution occurs everywhere !

## Slide 23: Digital transformation
Added information
Added information
V1.0
V2.0
V3.0
V4.0

## Slide 24: Digital transformation
Mechanical Measurements
Other measurement data
Third party
Post-processing
Many types of files, lot of exports to convert files and huge data storage;
Difficulties to merge data from several systems due to the time stamping, sampling frequency, meta data, labels, etc…
Combined data are not accessible real time.

## Slide 25: Digital transformation
Measurement
Input
Applications
Clients
Mechanical Measurements
Other measurement data
Third party
Client XY
Client Scope
Client Analysis
…
Data from every systems are continuously streamed and stored on the Cloud;
Data are open and accessible with several type of software's for monitoring and/or post-processing;
Real time access by several users asking different tasks;

## Slide 26: Digital transformation in conclusion
Digital transformation revolution occurs everywhere and open a new way to measure !
Industrial companies have already implemented this kind of solution based on Cloud technologies;
Cloud technology start to be available on several commercial measurement solutions :
Request some efforts to standardized our processes (labels, etc…) as it was done in the automotive industries 20 year’s ago with ASAM-ODS ;
For the mechanical measurements at CERN, the Cloud technology will be deploy this year on specific project (Test mode), ready to take part ?

## Slide 27: Thank you ! Questions ?