## Slide 1: FREIA Laboratory Facility for Research Instrumentation and Accelerator Development
Mitigation measures being studied and implemented at FREIA
Akira Miyazaki
for the FREIA Team
Burst Disc Review
21 May 2019

## Slide 2: Comparison of FREIA on April 9th and ESS
  Globally quenched Q0 matches Qext
    Inevitable due to physics!
The causes of the burst disc in FREIA on April 9th
  Self-Excited Loop (SEL)
   Locked the quenched cavity (kHz shift)
  MP conditioning at 4.2K  small margin (1.0 bar < 1.99 bar)
  Safety systems were developed
ESS machine (intrinsically safer)
  Generator driven + piezo tuner (400 Hz)
   the quenched cavity cannot be driven by RF
  Operation at 2.0K  safer (30mbar << 1.99 bar)
  Quench might cause some instability in the cryogenic plant  down time for beam?
Surface RRR=5
Qext
Surface RRR=10
Surface RRR=300
Prediction by BCS
Surface RRR=5
Surface RRR=10
Surface RRR=300

## Slide 3: Three options to consider
The risk of quench is inevitable in bulk Nb mitigation of the risks
Human error due to practical issues
  Two LLRFs: Lund system (pulse generator+ piezo compensation not deployed) and FREIA LLRF (Self-Excited Loop), combined circuit
  Two power stations: Electrosys (currently not working), DB (the only option)
  Three agilent power sensors for (Pf, Pr, Pt) (+ one extra) directly connected to LabView machine for SEL control
   Pragmatic decision: manually swapping all the component every time
Slow interlock from cryogenics
  Over pressure, LHe level, temperature after the quench
  At 2K (30mbar) rather than 4K
  Not implemented before the incident
Fast interlock by quench detection with LLRF
  Cut RF before or a.s.a.p. after the quench
  Quick solution in cryomodule testing at FREIA
  Possible impact in the beam operation  R&D interest in our LLRF system

## Slide 4: Three options to consider
The risk of quench is inevitable in bulk Nb mitigation of the risks
Human error due to practical issues
  Two LLRFs: Lund system (pulse generator+ piezo compensation not deployed) and FREIA LLRF (Self-Excited Loop), combined circuit
  Two power stations: Electrosys (currently not working), DB (the only option)
  Three agilent power sensors for (Pf, Pr, Pt) (+ one extra) directly connected to LabView machine for SEL control
   Pragmatic decision: manually swapping all the component every time
Slow interlock from cryogenics
  Over pressure, LHe level, temperature after the quench
  At 2K (30mbar) rather than 4K
  Not implemented before the incident
Fast interlock by quench detection with LLRF
  Cut RF before or a.s.a.p. after the quench
  Quick solution in cryomodule testing at FREIA
  Possible impact in the beam operation  R&D interest in our LLRF system

## Slide 5: Actions on the limited resource issue
Swapping the circuit was one of the causes of the incident
 We are working on improving our system
Duplicate the SEL circuit
  Implementing a new FPGA (new NI product)
  Two independent SELs will be ready before the 1st series cryomodule
Working on the Electrosys station
  One tube is found to be shorted
  Waiting for the arrival of the new tube
  Not ready for the prototype cryomodule testing
Duplicate power sensors
  Two extra power sensors + one power meter were ordered and received
  In total six power sensors + three power meter
  Enough to simultaneously monitor (Pf, Pr, Pt) of both cavities
  Deployment will be before the 1st series cryomodule
Work in progress but need to be pragmatic during prototype testing

## Slide 6: Reduce the possibility of human errors
Limited resources available
 recycling the circuit components and swapping CAV1 and CAV2 are unavoidable
Human errors on April 9th and our actions
Misconnection when the cable was swapped
Miscalculation of Eacc in the software
Check the software readout by sending RF by a signal generator from the bunker
  Expected Pf signal after the directional coupler (cable attenuation)
  Expected Pt signal from the cavity (cable attenuation + provided Qext, pu)
No visual monitor of all the important parameters
Monitor pressure, temperature, LHe level, radiation level, power levels, always visible during RF powering (from EPICS and power sensors)
  Emergency stop button
Lack of expert personal dedicated to the cryomodule testing
Involvement and training of other colleagues in FREIA
Check-list
The risk of human errors was drastically mitigated, but the system should be designed so that the safety does not depend on the operators

## Slide 7: Three options to consider
The risk of quench is inevitable in bulk Nb mitigation of the risks
Human error due to practical issues
  Two LLRFs: Lund system (pulse generator+ piezo compensation not deployed) and FREIA LLRF (Self-Excited Loop), combined circuit
  Two power stations: Electrosys (currently not working), DB (the only option)
  Three agilent power sensors for (Pf, Pr, Pt) (+ one extra) directly connected to LabView machine for SEL control
   Pragmatic decision: manually swapping all the component every time
Slow interlock from cryogenics
  Over pressure, LHe level, temperature after the quench
  At 2K (30mbar) rather than 4K
  Not implemented before the incident
Fast interlock by quench detection with LLRF
  Cut RF before or a.s.a.p. after the quench
  Quick solution in cryomodule testing at FREIA
  Possible impact in the beam operation  R&D interest in our LLRF system

## Slide 8: Interlock diagram on April 9th
LLRF
DB station
RF interlock
Arc1 det01
Arc1 det02
Arc1 det03
Arc4
det04
det05
det06
NI Compact RIO
TPG ch1
TPG ch2
e– pick ch1
e– pick ch2
PLC
Lund system
CAV2
Directional
coupler
Pf
CAV1 doorknob
CAV2 doorknob
Not used
CAV2 load
Test load
CAV2 circulator
CAV1
CAV2
CAV2 vacuum
CAV1 vacuum
Access
Water flow
Coupler T

## Slide 9: Interlock diagram on May 21th
LLRF
DB station
RF interlock
Arc1 det01
Arc1 det02
Arc1 det03
Arc4
det04
det05
det06
NI Compact RIO
TPG ch1
TPG ch2
e– pick ch1
e– pick ch2
PLC
Lund system
CAV2
Directional
coupler
Pf
CAV1 doorknob
CAV2 doorknob
Not used
CAV2 load
Test load
CAV2 circulator
CAV1
CAV2
CAV2 vacuum
CAV1 vacuum
LabView
EPICS
Cryogenic
PLC
MKS High
MKS Low
He vac. bar
He vac. mbar
Access
Water flow
Coupler T
LHe Level
Pt

## Slide 10: Interlock diagram on May 21th
LLRF
DB station
RF interlock
Arc1 det01
Arc1 det02
Arc1 det03
Arc4
det04
det05
det06
NI Compact RIO
TPG ch1
TPG ch2
e– pick ch1
e– pick ch2
PLC
Lund system
CAV2
Directional
coupler
Pf
CAV1 doorknob
CAV2 doorknob
Not used
CAV2 load
Test load
CAV2 circulator
CAV1
CAV2
CAV2 vacuum
CAV1 vacuum
LabView
EPICS
Cryogenic
PLC
MKS High
MKS Low
He vac. bar
He vac. mbar
Access
Water flow
Coupler T
LHe Level
Pt
MKS analog input split into two (most critical interlock)
Cryogenic regulation from MKS
Compact RIO’s analog input (100 kHz sampling)
If the pressure level is higher than 67 mbar, RF interlock trips
Pressurizing time (1.05 o 1.7 bar) after the quench was 20 s
 Enough safety margin from 30 mbar by this hardware interlock

## Slide 11: Interlock diagram on May 21th
LLRF
DB station
RF interlock
Arc1 det01
Arc1 det02
Arc1 det03
Arc4
det04
det05
det06
NI Compact RIO
TPG ch1
TPG ch2
e– pick ch1
e– pick ch2
PLC
Lund system
CAV2
Directional
coupler
Pf
CAV1 doorknob
CAV2 doorknob
Not used
CAV2 load
Test load
CAV2 circulator
CAV1
CAV2
CAV2 vacuum
CAV1 vacuum
LabView
EPICS
Cryogenic
PLC
MKS High
MKS Low
He vac. bar
He vac. mbar
Access
Water flow
Coupler T
LHe Level
Pt
Cryogenic PLC digital output
If LHe level becomes only a few cm above the cavity, RF interlock trips
Usually slower than the pressurizing
Redundancy in the safety system

## Slide 12: Interlock diagram on May 21th
LLRF
DB station
RF interlock
Arc1 det01
Arc1 det02
Arc1 det03
Arc4
det04
det05
det06
NI Compact RIO
TPG ch1
TPG ch2
e– pick ch1
e– pick ch2
PLC
Lund system
CAV2
Directional
coupler
Pf
CAV1 doorknob
CAV2 doorknob
Not used
CAV2 load
Test load
CAV2 circulator
CAV1
CAV2
CAV2 vacuum
CAV1 vacuum
LabView
EPICS
Cryogenic
PLC
MKS High
MKS Low
He vac. bar
He vac. mbar
Access
Water flow
Coupler T
LHe Level
Pt
Two software interlocks deployed
Interlock by He pressure
Cut LLRF’s DAC if He pressure reaches 100 mbar
Monitoring EPICS reading
He pressure < manually setup
Cavity temperature < 5K
Radiation < 7000 µSv/h
LHe < 55cm
Cavity vacuum 1e-6 mbar
Pf<, Pr<, Pt< manual setup
Cut HV of the power station
Again, redundancy in the safety system

## Slide 13: Three options to consider
The risk of quench is inevitable in bulk Nb mitigation of the risks
Human error due to practical issues
  Two LLRFs: Lund system (pulse generator+ piezo compensation not deployed) and FREIA LLRF (Self-Excited Loop), combined circuit
  Two power stations: Electrosys (currently not working), DB (the only option)
  Three agilent power sensors for (Pf, Pr, Pt) (+ one extra) directly connected to LabView machine for SEL control
   Pragmatic decision: manually swapping all the component every time
Slow interlock from cryogenics
  Over pressure, LHe level, temperature after the quench
  At 2K (30mbar) rather than 4K
  Not implemented before the incident
Fast interlock by quench detection with LLRF
  Cut RF before or a.s.a.p. after the quench
  Quick solution in cryomodule testing at FREIA
  Possible impact in the beam operation  R&D interest in our LLRF system

## Slide 14: Quench detection
For the cryomodule testing, it is enough to cut the RF after the quench
However, the quench event may cause instability in the cryogenic plant even if it is not destructive as bursting rupture disc
 Quench prevention before the event may be beneficial to reduce the machine down time in ESS
Experiences in Eu-FEL@DESY (and STF@KEK)
J. Branlard, et al., THPPC072, Proc. ICALEPCS2013, Sanfrancisco, USA
One quench  1 hour down time
More cavities  more down time
They deployed a quench detection system in LLRF and cut the RF beforehand.

## Slide 15: Quick solution by the existing system
LLRF
DB station
RF interlock
Arc1 det01
Arc1 det02
Arc1 det03
Arc4
det04
det05
det06
NI Compact RIO
TPG ch1
TPG ch2
e– pick ch1
e– pick ch2
PLC
Lund system
CAV2
Directional
coupler
Pf
CAV1 doorknob
CAV2 doorknob
Not used
CAV2 load
Test load
CAV2 circulator
CAV1
CAV2
CAV2 vacuum
CAV1 vacuum
LabView
EPICS
Cryogenic
PLC
MKS High
MKS Low
He vac. bar
He vac. mbar
Access
Water flow
Coupler T
LHe Level
Pt
Below THR fast interlock in Lund system
If transmitted waveform drops below THR, cut the RF through compact RIO
THR

## Slide 16: More sophisticated algorithm
Before the hard quench, usually an anomaly in the field decay is observed
  Partial quench  Q0 drop  QL drop  shorter time constant
STF’s THR: 1/2 QL Cut RF 2-3 pulses before the hard quench
System requirements
Fast RF digitizer + FPGA if FPGA resource is enough
DESY deploys this in a software running in a server, not in FPGA

## Slide 17: Our proposal
The field decay is
Differential operation
Numerical differential can provide in situ value during the decay
Pros: Required resources in the FPGA is smaller than fitting or logarithmic operation
Cons: Numerical differentiation is sensitive to the noise
Algorithm under development 
Required resource ~ only 5% of total FPGA slices
Deployment into ESS test stand/machine is to be discussed

## Slide 18: Summary
The burst disk incident on April 9th was investigated and the mitigation procedure was studied and implemented
  No identical risk to the ESS machine
The risk of human errors was mitigated by
  Double checking the read-out (even if it seems too redundant)
  Intuitive monitoring software
  Coherent work with different experts: RF, cryogenics, LLRF
  Working on duplicating all the circuit components to avoid swapping them (near future)
The slow interlock system was deployed to avoid the disc bursting
  Pressure interlock (hardware)
  Cryogenic interlock (hardware)
  Software interlock (even if it seems too redundant)
Fast quench prevention is not fully deployed but
  Below threshold interlock in the present Lund system is used as a first-aid
  Preliminary R&D project is ongoing to profit from QL anomaly originally developed in DESY and KEK
  Might be essential to the ESS machine to reduce the down time
  Algorithm is developed in FREIA LLRF  Deployment deserves ESS’s decision

## Slide 19: Appendix 1
Interlock from e- pickup
Digital output produced inside the e- pickup box
  THR defined by a potentiometer inside the box (2mA)
Analog read-out and interlock after cRIO’s ADC (100 kHz)
  Min and max THR adjustable by EPICS and are (0V, 1.5V)
NI Compact RIO
e– pick
Box
DO
DI
AI
AO

## Slide 20: Appendix 2
Interlock from beam vacuum
Digital output produced inside TPG
  THR defined at 1e-5 mbar
Analog read-out and interlock after cRIO’s ADC (100 kHz)
  Min and max THR adjustable by EPICS and are (0V, 2e-7 mbar)
NI Compact RIO
TPG300
DO
DI
AI
AO

## Slide 21: Appendix 3
Interlock from GHe pressure
Split analog reading and interlock after cRIO’s ADC (100 kHz)
  Min and max THR adjustable by EPICS and are (13.3, 67 mbar)
NI Compact RIO
MKS
AI
AO
Vacuum sensor
For cryogenic regulation
Two different dynamics ranges  two analog paths
Higher: 0-2666 mbar
Lower: 0-133 mbar