## Slide 1: FREIA Laboratory Facility for Research Instrumentation and Accelerator Development
Spoke Cryomodule Testing
Roger Ruber, for the FREIA team
Uppsala, 26 August 2019

## Slide 2
Overview of Infrastructure for Cryomodule Testing
FREIA infrastructure

## Slide 3: FREIA Laboratory
cryogenics
- liquid helium
- liquid nitrogen
control room
- equipment controls
- data acquisition
radio-frequency (RF) power sources
3 bunkers
with test stands
horizontal cryostat
vertical cryostat
State-of-the-art Equipment
Facility for Research Instrumentation and Accelerator Development
Funded by KAWS, Government, Uppsala Univ.
Competent and motivated staff
collaboration of physics (IFA) and engineering (Teknikum).

## Slide 4: Cryogenics
Helium liquefaction
  150 l/h at 4.5K (LN2 pre-cooling)
  2000 l LHe dewar/buffer, 3+1 outlets
  cryostats connected in closed loop
Gas recovery
  100 m3 gasbag
  3x 25 m3/h compressor
  10 m3 200 bar storage
2K Pumping
  ~3.2 g/s at 10 mbar
  ~4.3 g/s at 15 mbar
  110(90)W at 2.0(1.8)K
Liquid nitrogen
  20 m3 LN2 tank
Sub-atmospheric pumping station
Helium gas recovery system

## Slide 5: High Power RF Amplifiers
400 kW pulsed (352 MHz)
  2 stations, each 2 tetrodes TH595(A)
  3.5 ms, 14-28 Hz
50 kW CW (352/400 MHz)
  single tetrode TH571b
RF Switch Board (coax)
  ~15 min switch time

## Slide 6: FPGA Based Digital LLRF
Signal Driven
2 ADC inputs at 250 Msps
  (*) analogue bandwidth of 750 MHz
2 DAC outputs at 500 Msps
Digital downconversion to baseband 0 Hz, no analog mixers
  downconverted signal at 10 Msps or 1 Msps, selectable
undersampling to operate at any frequency from 10 to 750 MHz*
Self-excited Loop
CW or
pulsed mode
  switch closes the loop for a duration of 3.2 ms, repetition rate of 14 Hz.
Code in the FPGA

## Slide 7: Controls and Measurement
Slow Controls & Safety Interlock
  quench detection
  connecting different sub-systems: cryogenics, cryostats, powering, ...
  Siemens PLC + Nat.Instr. cRIO
  LabVIEW
    power coupler auto conditioning
      pulse length, power level, (rep rate=14 Hz)
    SEL control and data acquisition
    monitor cavity profile (IQ)
    cavity frequency shift tracing
    dynamic Lorentz force detuning
  EPICS
    user interface
    data archiver

## Slide 8: Ongoing Maintenance Work
Cryomodule
  water flow meters
    added filter, installed new meters → seems ok for the moment
Cryogenic system
  loosing small amount of helium gas during operation
    suspect leak in the recovery compressors
  recovery compressors run below capacity
    looking for replacement compressor
Control system
  implementing the valve box and cryomodule in EPICS
  LLRF implementation (new system version)
Improving software
  automatic conditioning and measurement software

## Slide 9
Prototype Spoke Cryomodule Testing
cryomodule Experience so far

## Slide 10: Spoke Valve Box Prototype
Cryo test run with simulator (Dec + Jan)
  commissioning controls & functionality test
Thermo-acoustic oscillations
  installed RLC-circuit to damp the oscillations

## Slide 11: Warm RF Conditioning
~3 days/cavity
  1st conditioning ~50 hours
  2nd conditioning ~20 h (cross-contamination)
fully automatic system conditioning system
  commissioned and implemented to run 24h/day
MP bands consistent with HNOSS test
  strength depends on pulse length,
  1st/2nd conditioning…

## Slide 12: Cool Down
Cool Down
  4h 15min to 4.5 K
    2h 30 min for shield
  ~1/2 day to 2 K
    including stabilization

## Slide 13: Cold RF Conditioning
FPC conditioning (4K)
  no coupler activity
  Quench during cavity conditioning at 4 K
    burst disc rupture → thermal cycling
    implemented quench detection
    from now on always at 2 K
FPC conditioning (2K)
  ~1/2 day/cavity
FPC1
Pulse length
vacuum
Forward power (dBm)
Pulse length
vacuum
Forward power (dBm)
FPC2

## Slide 14: Overview FPC Conditioning
Re-conditioning is required due to cross contamination
Total conditioning timing was
  about 4 weeks (real time),
  can be improved to 7 days (based on live time) or
  3 days if conditioning two couplers simultaneously
FPC1 | Warm | Cold
Real time [h] | 304 | 4.1
Intervention # | 17 | 2
Live time [h] | 76 | 2.3
Dead time [h] | 228 | 1.8
FPC2 | Warm | Cold
Real time [h] | 297 | 2.9
Intervention # | 23 | 1
Live time [h] | 78 | 2.8
Dead time [h] | 219 | 0.1

## Slide 15: Cavity Conditioning
Different loop setting and conditioning strategies were studied
  no significant difference was found in the time variation of cavity conditioning.
Open loop with a fixed repetition rate of 14 Hz would be an optimal option
Total cavity conditioning timing is about 2 days
 | Cav1 | Cav2 1st run | Cav2 2nd run | Cav2 3rd run
live time [h] | 6 | 6.7 | 7.3 | 7.4
Loop | Open loop | SEL | SEL | SEL
Conditioning strategy | 1ms@ 14Hz 3.2ms@14Hz | 1ms@ 14Hz 3.2ms@14Hz | 3.2ms@ 1Hz 3.2ms@14Hz | 1ms@ 14Hz 3.2ms@14Hz

## Slide 16: Spoke Cryomodule Prototype (2019)
Fast tuner performance
  Lorenz force detuning compensation (piezo)
w/o correction
w/ correction
fshift
Pfwd
Pref
Ptrans

## Slide 17: Spoke Cryomodule Prototype (2019)
Cavity #2 performance
  multipacting regions similar as prototype
    2-3; 4-5; 7-8 MV/m
  field emission sensitive to tuner motion or position (under investigation)
  thermal cycle cryomodule up to 40 K
    outgassing up to 10-4 mbar observed around 20 K,
      implying that hydrogen is vaporized from the cavity surface.
    field emission almost back as before the CTS activities
  maximum Eacc improved from 9 MV/m to 10 MV/m after thermal cycle
  Q0 performance of cavity2 recovered as before the CTS activities
FREIA CAV2 after tuner test (2019)

## Slide 18
Series Spoke Cryomodule Testing
how to keep testing to 4 weeks, and
which additional resources required to achieve that?
outlook

## Slide 19: There Will Be Differences ...
Prototype to 1st series cryomodule
  5 people from IPNO did the prototype installation
    none of FREIA was unfortunately able to follow up all the details
  we should learn when removing the prototype
  support frame / trolley is different
  doorknob is different
  instrumentation connectors are different
    have to redo connectors on local cabling (and commission)
1st to following series cryomodules
  pumping ports
    trade-off between speed and safety
    1st cryomodule can/will be tested with 2 pumps

## Slide 20: Testing and Operation
2 hours for moving the tuner (1 hour for one direction)
  thus there is an advantage to prolong the daily operation time
need & accuracy of the Q0 value
  measurement time depends on required accuracy
  is it enough to give the upper limit of the power dissipation?
RF stations reliability and maintenance is a problem
  need a good high power RF engineer
    previous announcements did not result in a good candidate → try again
LLRF: waiting for a new system (replacing an old version)
  hardware available, but no working system yet
    integrated control system part needs more work
Documentation and test results
  report should be ready before completing test next module/installation at ESS
  also: agree on what & how to test (on reception, cold test, ...)

## Slide 21: Critical Personnel Resources
During installation
  1st cryomodule:
    Patxi and Matthieu from IPNO, to teach us and check our work
  1st + following
    1 or 2 mechanics, to help us, and (if from ESS) for their training
During testing
  on-site: 3 operators in CR for cryo, SRF, LLRF/controls
    1 or 2 operators, to off-load Rocio/Lars + help extend daily operation time
  on-call: high power RF
    1 engineer, to assist Rolf & co
  remote: LLRF
    ICS/uTCA-LLRF expert from ESS (have no expertise at FREIA)

## Slide 22: Critical Spare Parts
RF stations
  extra RF station as back up, and/or
  preferably same amount of spare tetrode tubes as in use
    2 ordered: 1 to arrive end August, 1 in Feb 2020 (10 months delivery time!)
    order more from Thales and/or can we rely on ESS stock?
LLRF
  spare FPGA cards for Uppsala NI/LabVIEW system
  spare µTCA cards in Uppsala, as delivery from Lund takes a few days
Other
  cooling water flow meters (FPC)
    preparing Eletta flow switches as back-up
  RF power meters
    want to keep the ESS power meters to speed up calibration procedures

## Slide 23: Summary
Uppsala University & FREIA Laboratory doing their best efforts to make this project a success and on-time
Most critical resources
operator/technician
  mechanic during installation
  operator during testing
  ICS remote support for LLRF
high power RF engineer
spare parts
  tetrodes
  LLRF
acceptance criteria