## Slide 1: FREIA Laboratory Facility for Research Instrumentation and Accelerator Development Department of Physics and Astronomy
Projekt kalla magneter Möte om tillverkningsteknik
Roger Ruber
Uppsala, 3 March 2021

## Slide 2: Nuvarende MCBC/MCBY Orbit Corrector
..
Parameters | MCBC | MCBY
Coil inner diameter | 56.4 mm | 70.4 mm
Magnetic length | 0.904 m | 0.899 m
Operating temperature | 1.9 K and 4.5 K | 1.9 K and 4.5 K
Nominal field (at 1.9/4.5 K) | 3.11/2.33 T | 3.00/2.5 T
Nominal current (at 1.9/4.5 K) | 100/74 A | 88/72 A
Peak field in coil (at 1.9/4.5 K) | 3.65/2.68 T | 3.60/2.96 T
Short sample current (at 1.9/4.2 K) | 172/127 A | 162/120 A
Stored energy (at 1.9 K) | 14.2 kJ | 13.6 kJ
Self-inductance | 2.84 H | 5.27 H
DC resistance (RT) | 375 Ω | 501 Ω
Overall length | 1100 mm | 1100 mm
Outer diameter of assembly | 452 mm | 475 mm
Magnet module mass | 203 kg | 193 kg
Total mass including the support structure | 1234 kg | 1247 kg
Superconductor type | Nb-Ti in Cu matrix | Nb-Ti in Cu matrix
SC wire dimension | 0.38 mm x 0.73 mm (+0.0/-0.01) | 0.38 mm x 0.73 mm (+0.0/-0.01)
Ribbon construction | 14 wire (glued) | 15 wire (glued)

## Slide 3: Magnetdesign
Diskussion om högre ström
  begränsas av teknisk installation i LHC tunneln
  200 av denna typ av magnet
Power converter
  nuvarande ström circuit ±120 A, ± 10 V
    interlock för quench skydd:
      extern signal eller ström fel
    ingen energy extraction
  dokumentation
    http://sy-dep-epc-lpc.web.cern.ch/converters/lhc120a-10v/general.stm
    https://indico.cern.ch/event/904900/
Behöver passiv quench skydd
  diod och motstånd
  kall system, dvs inbyggd i magnetens kryostat
  att diskutera med CERN
FGC=function generator/controller

## Slide 4: Mekanisk design
plats för elektriska skarv (joint-boxes)
  50cm modell: joint-box vid varje end
  alternativ ide: fram/baksida av järn ok

## Slide 5: Elektrisk diagram (anpassas för repkabel)

## Slide 6: Supraledare
Bruker Energy & Supercon Technologies (BEST)
  https://www.bruker.com/content/bruker/int/en/products-and-solutions/superconductors.html
  supraledande tråd
    0.360 mm ± 0.006 mm
      bare 0.330 mm
      varnish (formvar) insulation
    Icritical (4.2K, 3T) ≥ 137 A
      at 60% load line = 82 A
    Copper to Superconductor ratio 1.35 : 1
    RRR 70 or 100
  kabel: 6-around-1
    förväntar 1.08 mm diameter
    60% load line current = 574 A
    6 ledare vridas kring central ledare som blir därmed 10% kortare än de andra 6
    fråga: maximum böjningsradi?

## Slide 7: Supraledare
längd i magnetspindel
  100 m för 1 winding
  10 windings per spår
  7 tråd per kabel
per magnetspindel
  100 x 10 x 7 m = 7 km tråd
  100 x 10 = 1 km rep-kabel
per magnet (= 2 spindel)
  2 km rep-kabel
CCT prototyp
  test spindel + slask
  2 magnet = 4 km
  1 reservmagnet = 2 km
  TOTAL = 10 km rep-kabel

## Slide 8: Magnetspindel
A6082-T6
  yield strength 400 MPa
  notch-yield ratio = tear resistance to yield strength
  operate below 120 MPa
  high RRR (quench skydd)
anodisering
  ändrar elektrisk motstånd
    ger extra isolation för kabeln
  hjälper anknytning av epoxy resin
  för 50cm modell magnet
    40 µm tjock
3 olika spindel
  2 med spår och 1 yttre stöd-spindel
  3D step modell transfer från CAD till CNC
viktig med dimension kontroll
Ratio > 1.5
AL 6082-T6
50cm CCT modell
Oxford University 2018

## Slide 9: Magnetspindel
Erfarenhet Glyn
  spår max. 5 djup x 1 bred
Design
  spår pitch 3.1 mm
  bred 2 x 1.08 + 0.1 = 2.3 mm
  djup 5 x 1.08 + 0.1 = 5.5 mm
50cm CCT modell
Min wall 0.3mm with 2mm wide channel

## Slide 10: Magnetspindel
Spår i magnetspindeln
  7 wire per assemblage: ~1.1 mm
    spår blir ~2.3 mm x 6 mm

## Slide 11: Järn ok
Förslag för järn typ: ARMCO
  https://alliedmet.com/alliedpureiron/
diskussion med CERN
  this product is basically good. ARMCO or PURE IRON is the same. 5 mm or 6 mm thickness plates are standard values for the market. Originally CERN was using MAGNETIL,
  In the ... production for HL-LHC (5.8 mm, old LHC standard thickness) we have adopted the ARMCO Grade 4 (not the ARMCO grade 2) with a lower content of S, N, et (the presence of these components affecting the magnetic behaviour).
  ... had problems also with scaling (at the surface of the plates) and with the content of S (sometimes higher than 0.003, in this case we rejected the plates/heat).
  The supplier had problems in obtaining the correct grain size number too. We could never obtain the target value of Gs<4 (in the MS at the beginning CERN asked fro Gs< 3!). The Gs is important for the mechanical characteristics at low temp. Finally, the annealing process included in the spec in order to favourite the Gs reduction had to be abandoned.

## Slide 12: Järn ok
MCBC/Y design har stor öppning för magneten
  CCT magnet design är mindre
  mer plats för järn
    förbättring av fältkvalitet
    mindre ström
MCBC/Y
alternative
pin
tie rod
cooling
keys
alignment slot

## Slide 13
1...5.8±0.15
pin & rod
4x keys
guide thermal contraction