## Slide 1: Projekt kalla magneter
Tillverkningsteknik
Roger Ruber för projektgruppen
Uppsala, 14 April 2021

## Slide 2: Nuvarande MCBC/MCBY Orbit Corrector
Kombination
B0 = 3.11 T
Lmagnetic = 0.904 m
= 2.81 Tm
aperture = 35.2 mm
Imax = 100 A
Parameters | MCBC | MCBY
Coil inner diameter | 56.4 mm | 70.4 mm
Magnetic length | 0.904 m | 0.899 m
Operating temperature | 1.9 K and 4.5 K | 1.9 K and 4.5 K
Nominal field (at 1.9/4.5 K) | 3.11/2.33 T | 3.00/2.5 T
Nominal current (at 1.9/4.5 K) | 100/74 A | 88/72 A
Peak field in coil (at 1.9/4.5 K) | 3.65/2.68 T | 3.60/2.96 T
Short sample current (at 1.9/4.2 K) | 172/127 A | 162/120 A
Stored energy (at 1.9 K) | 14.2 kJ | 13.6 kJ
Self-inductance | 2.84 H | 5.27 H
DC resistance (RT) | 375 Ω | 501 Ω
Overall length | 1100 mm | 1100 mm
Outer diameter of assembly | 452 mm | 475 mm
Magnet module mass | 203 kg | 193 kg
Total mass including the support structure | 1234 kg | 1247 kg
Superconductor type | Nb-Ti in Cu matrix | Nb-Ti in Cu matrix
SC wire dimension | 0.38 mm x 0.73 mm (+0.0/-0.01) | 0.38 mm x 0.73 mm (+0.0/-0.01)
Ribbon construction | 14 wire (glued) | 15 wire (glued)

## Slide 3: Elektrisk strömstyrka
Maximum ström
  begränsas av teknisk installation i LHC tunneln
  200 av denna typ av magnet
Power converter
  nuvarande ström circuit ±120 A, ± 10 V
    interlock för quench skydd:
      extern signal eller ström fel
    ingen energy extraction
  dokumentation
    http://sy-dep-epc-lpc.web.cern.ch/converters/lhc120a-10v/general.stm
    https://indico.cern.ch/event/904900/
Behöver passiv quench skydd
  diod och motstånd
  kall system, dvs inbyggd i magnetens kryostat
  att diskutera med CERN
FGC=function generator/controller

## Slide 4: Supraledare - tråd
Bruker Energy & Supercon Technologies (BEST)
  https://www.bruker.com/content/bruker/int/en/products-and-solutions/superconductors.html
supraledande tråd
  diameter
    0.33 mm (w/o insulation = bare)
    0.36 mm (with insulation)
    insulation = formvar (or polyimide)
  filament approx. 29 µm
  bending radius 25 mm
  max. tensile strength 20 kg/mm2
  Icritical (4.2K, 3T) ≥ 137 A
    at 60% load line = 82 A
  koppar till supraledare förhållandet 1.35 : 1
  RRR ≥70

## Slide 5: Supraledare - tråd
Bruker has send 100 m of wire
  shipment to Uppsala University (Kevin)
5 m to UU
  testing joints for installation in cryostat
5 m to CERN
  Ic testing between 2 and 5 T
    1.6 m per sample
  cross section investigation (photo)
remainder to Scanditronix
  test making joints
  make 6-around-1 cable
    check diameter & bending radius
    impregnation into channel (100 to 200 mm)
      check impregnation
      test insulation (let wire stick out of channel)
Cable arrived today 14-Apr

## Slide 6: Superconductor - Quench Integral
from Glyn

## Slide 7: Supraledare - kabel
  kabel: 6-around-1
    förväntat diameter 1.08 mm
    Icritical (4.2K, 3T) ≥ 7x137 A = 959 A
      60% load line current = 574 A
    men koppla alla 7 ledare i serie → samma Icritical som tråden
  6 ledare vridas kring central ledare som blir därmed 10% kortare än de andra 6
    förväntas inte vara ett problem
  fråga: maximum böjnings radie?
    böjnings radie tråd = 25 mm
1.08 mm

## Slide 8: Superconducting Cable – Questions (1/2)
From Glyn (12-Apr-2021):
there must be a formula for the twist pitch that is possible and the resulting dia of the rope.
should we twist CW or ACW?
what is the difference in length of the spiraling outer strands to the central strand.
  Bruker mentioned ~10% with their cable twist
what is the effect on inductance. the strands in the rope must have a small increase in inductance! is this important or not. we could measure when we have it may be?
stress in a twisted rope when the magnetic field / hoop stress is applied? it must be a small effect but we should think about it.
as a thought experiment , imagine the twisted rope, wind it into the channel. the wires must slide relative to each other a little. if an axial tension is applied to the rope the outer conductors must try to squeeze the central conductor. we need to think about the insulation point stresses in the rope. there must be some papers on this?

## Slide 9: Superconducting Cable – Questions (2/2)
it a rather deep in consideration and we can probably ignore it but someone will ask. don't want the insulation to fail!
to keep rope from springing open, do we need to put an outer tape on it? may be? this will change the channel width.
i was thinking we should wind the rope, open the wires at the ends, and check the insulation between wires! so need to think about how to fix the ends without shorts.
we could put the ropes into the channel, impregnate with some resign or wax! press the ropes into the channel with a small bar at the pressure we expect to see in the magnet and check the insulation between wires.?
what is the effect of the twist on the current, it has some effect on the effective current due to the twist angle.
what is the field in the cable, just a variation of engineering current density?

## Slide 10: Supraledare - längd
spår i magnetspindel
  166 vändningar
  2x5 = 10 lindningar per spår
längd i magnetspindel
  64,85 m per lindning
  648,5 m total kabel längd
  7 tråd per kabel
    7 x 64,85 m = 453,95 m tråd per lindning
    7 x 648,5 m = 4539,5 m tråd total
längd i magnet (2 magnetspindel)
  129,7 m per lindning
  1297 m total kabel längd
  7 tråd per kabel
    7 x 129,7 m = 908 m tråd per lindning
    7 x 1297 m = 9079 m tråd
per magnetspindel
  min. längd 65 m + extra = 70 m
  behövs 700 m rep-kabel
per magnet (= 2 spindel)
  min. längd 130 m + extra = 140 m
  behövs 1,4 km rep-kabel
CCT prototyp
  test spindel = 0,7 km
  2 magnet = 2,8 km
  1 reservmagnet = 1,4 km
  TOTAL = 5 km rep-kabel (= 35 km tråd)
    min. längd 140 m

## Slide 11: Elektrisk diagram (anpassas för repkabel)

## Slide 12: Magnetisk design
Winding pitch 5.5 mm
Mid-plane gap inner coil = 0.35 * cos (30 deg) = 0.303 mm
  166 turns corresponds to magnetic length of 913 mm and a total coil length of 1064 mm
I = 7*79A
  57.7% of short sample
B0 = 3.06 T
= 2.8 Tm

## Slide 13: Magnetisk belastning
aluminium A6082-T6
  elasticity = 70 GPa
  tensile strength 300 MPa
  yield strength 400 MPa = 4x108 N/m2

## Slide 14: Mekanisk design
plats för elektriska skarv (joint-boxes)
  50cm modell: joint-box vid varje end
  alternativ ide: fram/baksida av järn ok

## Slide 15: Magnetspindel
A6082-T6
  yield strength 400 MPa
  notch-yield ratio = tear resistance to yield strength
  operate below 120 MPa
  high RRR (quench skydd)
anodisering
  ändrar elektrisk motstånd
    ger extra isolation för kabeln
  hjälper anknytning av epoxy resin
  för 50cm modell magnet
    40 µm tjock
3 olika spindel
  2 med spår och 1 yttre stöd-spindel
  3D step modell transfer från CAD till CNC
viktig med dimension kontroll
Ratio > 1.5
AL 6082-T6
50cm CCT modell
Oxford University 2018

## Slide 16: Magnetspindel
Erfarenhet Glyn
  spår max. 5 djup x 1 bred
Design
  spår pitch 3.1 mm
  bred 2 x 1.08 + 0.1 = 2.3 mm
  djup 5 x 1.08 + 0.1 = 5.5 mm
50cm CCT modell
Min wall 0.3mm with 2mm wide channel

## Slide 17: Magnetspindel
Spår i magnetspindeln
  7 wire per assemblage: ~1.1 mm
    spår blir ~2.3 mm x 6 mm

## Slide 18: Järn ok
Förslag för järn typ: ARMCO
  https://alliedmet.com/alliedpureiron/
diskussion med CERN
  this product is basically good. ARMCO or PURE IRON is the same. 5 mm or 6 mm thickness plates are standard values for the market. Originally CERN was using MAGNETIL,
  In the ... production for HL-LHC (5.8 mm, old LHC standard thickness) we have adopted the ARMCO Grade 4 (not the ARMCO grade 2) with a lower content of S, N, et (the presence of these components affecting the magnetic behaviour).
  ... had problems also with scaling (at the surface of the plates) and with the content of S (sometimes higher than 0.003, in this case we rejected the plates/heat).
  The supplier had problems in obtaining the correct grain size number too. We could never obtain the target value of Gs<4 (in the MS at the beginning CERN asked fro Gs< 3!). The Gs is important for the mechanical characteristics at low temp. Finally, the annealing process included in the spec in order to favourite the Gs reduction had to be abandoned.

## Slide 19: Järn ok
MCBC/Y design har stor öppning för magneten
  CCT magnet design är mindre
  mer plats för järn
    förbättring av fältkvalitet
    mindre ström
MCBC/Y
alternative
pin
tie rod
cooling
keys
alignment slot

## Slide 20
1...5.8±0.15
pin & rod
4x keys
guide thermal contraction

## Slide 21: Next Steps
discussing collaboration contract with CERN
  to receive full 3D CAD model from CERN
  to receive advise/knowledge (Glyn Kirby) and engineering support
recruiting researcher/post-doc for simulations
  work on field and quench simulations, magnetic design
  discussion ongoing with candidate presently at CERN
tests on superconducting cable
  joints, cabling, impregnation
tests cutting iron for yoke laminations
  accuracy
submitted abstract to MT27 Magnet Technology Conference
  in Japan, 15-19 November 2021